## 3186 PLC 03-25-11 JB
outputs
O:0/0 | BLWR_EN | Digital | 24VDC | Switched | DIGITAL OUTPUT: | 9
O:0/1 | BLWR_RST | Digital | 24VDC | Switched | ANALOG OUTPUT: | 7
O:0/2 | PUMP_EN | Digital | 24VDC | Switched
O:0/3 | PUMP_RST | Digital | 24VDC | Switched | ANALOG INPUT: | 10
O:0/4 | AIR_HEAT_CON | Digital | 120
O:0/5 | GLY_HEAT_CON | Digital | 120
O:0/8 | COMP_CON | Digital | 120
O:0/9 | HG_SOL | Digital | 120
O:0/10 | SC_SOL | Digital | 120
O:1.0 | BLOWER_SPEED | Analog 4-20mA
O:1.1 | PUMP_SPEED | Analog 4-20mA
O:1.2 | BYPASS_DMPR | Analog 4-20mA
O:1.3 | AIR_EXV | Analog 4-20mA
O:1.4 | GLY_EXV | Analog 4-20mA
O:1.5 | AIR_SCR | Analog 4-20mA
O:1.6 | GLY_SCR | Analog 4-20mA
inputs
I:0/0 | GLY_TEMP_LSL | 120VAC | 24VDC
I:0/1 | GLY_HEAT_OT | 120VAC | 24VDC
I:0/2 | AIR_HEAT_OT | 120VAC | 24VDC
I:0/3 | BLOWER_ON | 120VAC | 24VDC
I:0/4 | PUMP_ON | 120VAC | 24VDC
I:0/5 | COMP_OL | 120VAC | 24VDC
I:0/6 | COMP_DIS_PSH | 120VAC | 24VDC
I:0/9 | EMO_CLR | 120VAC | 120VAC | Can complete with SS relays
I:0/10 | RESET_PB | 120VAC | 120VAC | NO relay needed
I:0/11 | MASTER_SW | 120VAC | 120VAC | Can complete with SS relays
bits from HMI
B32:0/0 | ALARM_RESET
B32:0/1 | HEARTBEAT | PLC sets high
B32:0/2 | HMI_FLT | shuts down equipment
B32:2/0 | AIR_SYSTEMS_EN
B32:2/2 | GLY_SYSTEMS_EN
bits to HMI
B33:2/0 | AIR_TEMP_ON
B33:2/1 | AIR_FLOW_ON
B33:2/2 | GLY_TEMP_ON
B33:2/3 | GLY_FLOW_ON
B33:5/0 | EMO_FLT
B33:5/6 | GLY_LSL_FLT
B33:5/8 | GLY_PRESS_FLT
B33:5/9 | AIR_HEAT_OT_FLT
B33:5/10 | BLOWER_FLT
B33:5/11 | PUMP_FLT
B33:5/12 | GLY_HEAT_OT_FLT
B33:5/13 | DIS_PRESS_FLT
B33:5/14 | COMP_OL_FLT
B33:5/15 | HEARTBEAT_FLT
integers from HMI
N30:30 | AIR_TEMP_HEAT_PCT
N30:31 | AIR_FLOW_BLOWER_PCT
N30:32 | GLY_TEMP_HEAT_PCT
N30:33 | GLY_FLOW_PUMP_PCT
N30:40 | AIR_TEMP_COOL_PCT
N30:41 | AIR_FLOW_DAMPER_PCT
N30:42 | GLY_TEMP_COOL_PCT
integers to HMI
N31:50 | PROD_IN_TEMP | vaisala
N31:51 | PROD_IN_HUM
N31:52 | NOZ_IN_TEMP
N31:53 | NOZ_IN_HUM
N31:54 | SUCT_PRESS_PSIA_X10 | must be psia x10
N31:55 | SPARE_2_5
N31:56 | SUCT_TEMP
N31:57 | DISCH_TEMP
N31:58 | BARO_PRESS_PSI_X1000
N31:59 | NOZ_DP_X1000
N31:60 | PROD_AIR_DP_XE5
N31:61 | PROD_GLY_DP_X1000
N31:62 | GLY_DENSITY_X100
N31:63 | GLY_FLOW_X10
N31:64 | GLY_PRESS_IN_X10
N31:65 | GLY_PRESS_OUT_X10
N31:66 | GLY_IN_1_TEMP
N31:67 | GLY_IN_2_TEMP
N31:68 | GLY_OUT_1_TEMP
N31:69 | GLY_OUT_2_TEMP
N31:70 | AIR_IN_1_TEMP
N31:71 | AIR_IN_2_TEMP
N31:72 | AIR_IN_3_TEMP
N31:73 | AIR_IN_4_TEMP
N31:74 | AIR_OUT_1_TEMP
N31:75 | AIR_OUT_2_TEMP
N31:76 | AIR_OUT_3_TEMP
N31:77 | AIR_OUT_4_TEMP
N31:78 | AIR_OUT_5_TEMP
N31:79 | AIR_OUT_6_TEMP
N31:80 | AIR_OUT_7_TEMP
N31:81 | AIR_OUT_8_TEMP

## NI USB-6259
OUTPUTS
PLC |  | WIRE | DESC | SIGNAL | CHASSIS | Contact | RELAY | COMMENTS
O:0/0 | BLWR_EN | 0 | Air blower RUN | 24VDC | USB-6259/P0.0 | 65 | 5VDC to 24VDC | Condensed into a single output via one relay, 24VDC sourced from AB VSD
O:0/1 | BLWR_RST | 1 | Air blower STOP | 24VDC
O:0/2 | PUMP_EN | 2 | Gly Pump RUN | 24VDC | USB-6259/P0.1 | 66 | 5VDC to 24VDC | Condensed into a single output via one relay, 24VDC sourced from AB VSD
O:0/3 | PUMP_RST | 3 | Gly Pump STOP | 24VDC
O:0/4 | AIR_HEAT_CON | 4 | Air heater CON | 120VAC | USB-6259/P0.2 | 67 | 5VDC to 120VAC
O:0/5 | GLY_HEAT_CON | 5 | Gly heater CON | 120VAC | USB-6259/P0.3 | 68 | 5VDC to 120VAC
O:0/8 | COMP_CON | 8 | Compressor CON | 120VAC | USB-6259/P0.4 | 69 | 5VDC to 120VAC | Incorrect label on DAQ side (currently as 00008)
O:0/9 | HG_SOL | 9 | Hot gas solenoid | 120VAC | USB-6259/P0.5 | 70 | 5VDC to 120VAC
O:0/10 | SC_SOL | 10 | Suction cool solenoid | 120VAC | USB-6259/P0.6 | 71 | 5VDC to 120VAC
O:1.0 | BLOWER_SPEED | 150 | Blower speed control | 4-20mA
O:1.1 | PUMP_SPEED | 151 | Pump speed control | 4-20mA
O:1.2 | BYPASS_DMPR | 152 | Bypass damper | 4-20mA
O:1.3 | AIR_EXV | 153 | Ref expansion valve AIR evap | 4-20mA
O:1.4 | GLY_EXV | 154 | Ref expansion valve GLY evap | 4-20mA
O:1.5 | AIR_SCR | 155 | air heater SCR | 4-20mA
O:1.6 | GLY_SCR | 156 | glycol heater SCR | 4-20mA
| OUTPUTS: |  | INPUTS:
| ON-OFF 24VDC | 4 | IN 24VDC | 8 | NI 9211 | MOD5
| ON-OFF 120VAC | 5 | IN 120VAC | 2
| ANALOG 4-20mA | 7 | 4-20mA | 11
|  |  | T-type TCs | 2
|  |  | frequency | 1
INPUTS:
PLC |  | WIRE | DESC | SIGNAL | CHASSIS | Contact | RELAY | COMMENTS
I:0/0 | GLY_TEMP_LSL | 10000 | GLY tank low level switch | 24VDC
I:0/1 | GLY_HEAT_OT | 10001 | GLY heater over temp | 24VDC |  |  |  | From High limit control
I:0/2 | AIR_HEAT_OT | 10002 | AIR heater over temp | 24VDC |  |  |  | From High limit control
I:0/3 | BLOWER_ON | 10003 | Blower on | 24VDC |  |  |  | From VSD
I:0/4 | PUMP_ON | 10004 | Pump on | 24VDC |  |  |  | From VSD
I:0/5 | COMP_OL | 10005 | Compressor over limit | 24VDC
I:0/6 | COMP_DIS_PSH | 10006 | Compressor Discharge pressure | 24VDC
I:0/9 | EMO_CLR | 10009 | Emergency OFF button clear | 120VAC |  |  | 120VAC to 5VDC
I:0/10 | RESET_PB | 10010 | Base reset after EMO | 24VDC
I:0/11 | MASTER_SW | 10011 | Master power switch | 120VAC |  |  | 120VAC to 5VDC
SENSORS:
PLC |  | WIRE | DESC | SIGNAL | CHASSIS | Contact | RELAY | COMMENTS
I:2.0 | RH product temp | 10200 | Temperature -40 to 176 °F | 4 - 20mA | cDAQ-9174 / 9208 | 0 |  | 251 is reference
I:2.1 | RH product | 10201 | HUM 0~100% | 4 - 20mA | cDAQ-9174 / 9208 | 1 |  | 251 is reference
I:2.2 | RH nozzle temp | 10202 | Temperature -40 to 176 °F | 4 - 20mA | cDAQ-9174 / 9208 | 2 |  | 251 is reference
I:2.3 | RH nozzle | 10203 | HUM 0~100% | 4 - 20mA | cDAQ-9174 / 9208 | 3 |  | 251 is reference
I:2.4 | Ref suct pressure | 10204 | 0-200PSIG | 4 - 20mA | cDAQ-9174 / 9208 | 4 |  | 251 is reference
I:2.5 | SPARE | 10205 | SPARE | 4 - 20mA | cDAQ-9174 / 9208 | 5 |  | 251 is reference
I:2.6 | Comp suct temp | 301TE | Compressor suction temp | thermocouple | cDAQ-9174 / 9211 | 1 |  | T-Type thermocouple
I:2.7 | Comp disch temp | 302TE | Compressor discharge temp | thermocouple | cDAQ-9174 / 9211 | 2 |  | T-Type thermocouple
I:3.0 | Baro Pressure | 10300 | PT302 12.777 ~ 15.717 psia | 4 - 20mA | cDAQ-9174 / 9208 | 6 |  | 251 is reference
I:3.1 | Nozzle dP | 10301 | Nozzle dP 0 to 5 inWC | 4 - 20mA | cDAQ-9174 / 9208 | 7 |  | 251 is reference
I:3.2 | Air dP | 10302 | Air dP 0 to 0.1 inWC | 4 - 20mA | cDAQ-9174 / 9208 | 8 |  | 251 is reference
I:3.3 | Glycol dP | 10303 | GLY dP 0 to 25psid | 4 - 20mA | cDAQ-9174 / 9208 | 9 |  | 251 is reference
I:3.4 | GLY density | 10304 | frequency measurement
I:3.5 | GLY flow | 10305 | Mass flow 0 to 70lbm/min | 4 - 20mA | cDAQ-9174 / 9208 | 10 |  | 10355 is reference
I:3.6 | GLY inlet Pressure | 10306 | Product inlet Pressure 0-200psig | 1~6VDC |  |  |  | 251 is reference
I:3.7 | GLY outlet Pressure | 10307 | Product inlet Pressure 0-100psig | 1~6VDC |  |  |  | 251 is reference
Note on E2: Carlos Herrera: Updated voltage once all is changed to REV C dwg status
Note on E27: Carlos Herrera: Updated voltage once all is changed to REV C dwg status
Note on E40: Carlos Herrera: Updated voltage once all is changed to REV C dwg status

## cDAQ-9174
OUTPUTS
PLC |  | WIRE | DESC | SIGNAL | cDAQ-9174 (used..left) | Contact | RELAY | COUNTS (this kind.. Total) | COMMENTS
O:0/0 | BLWR_EN | 0 | Air blower RUN | 24VDC | NI9403 (1..31) | DIO.0 | 70-ODC5 / 5VDC ctrl .. 3-60VDC load | 1..1 | Condensed into a single output via one relay, 24VDC sourced from AB VSD
O:0/1 | BLWR_RST | 1 | Air blower STOP | 24VDC
O:0/2 | PUMP_EN | 2 | Gly Pump RUN | 24VDC | NI9403 (2..30) | DIO.1 | 70-ODC5 / 5VDC ctrl .. 3-60VDC load | 2..2 | Condensed into a single output via one relay, 24VDC sourced from AB VSD
O:0/3 | PUMP_RST | 3 | Gly Pump STOP | 24VDC
O:0/4 | AIR_HEAT_CON | 4 | Air heater CON | 120VAC | NI9403 (3..29) | DIO.2 | OAC5 (5VDC to 120VAC) | 1..3
O:0/5 | GLY_HEAT_CON | 5 | Gly heater CON | 120VAC | NI9403 (4..28) | DIO.3 | OAC5 (5VDC to 120VAC) | 2..4
O:0/8 | COMP_CON | 8 | Compressor CON | 120VAC | NI9403 (5..27) | DIO.4 | OAC5 (5VDC to 120VAC) | 3..5 | Incorrect label on DAQ side (currently as 00006), OK on relay side (00008)
O:0/9 | HG_SOL | 9 | Hot gas solenoid | 120VAC | NI9403 (6..26) | DIO.5 | OAC5 (5VDC to 120VAC) | 4..6
O:0/10 | SC_SOL | 10 | Suction cool solenoid | 120VAC | NI9403 (7..25) | DIO.6 | OAC5 (5VDC to 120VAC) | 5..7
O:1.0 | BLOWER_SPEED | 150 | Blower speed control | 4-20mA | NI9265_1 (1..3) | AO0-COM0 | N/A
O:1.1 | PUMP_SPEED | 151 | Pump speed control | 4-20mA | NI9265_1 (2..2) | AO1-COM1 | N/A
O:1.2 | BYPASS_DMPR | 152 | Bypass damper | 4-20mA | NI9265_1 (3..1) | AO2-COM2 | N/A
O:1.3 | AIR_EXV | 153 | Ref expansion valve AIR evap | 4-20mA | NI9265_1 (4..0) | AO3-COM3 | N/A
O:1.4 | GLY_EXV | 154 | Ref expansion valve GLY evap | 4-20mA | NI9265_2 (1..3) | AO0-COM0 | N/A
O:1.5 | AIR_SCR | 155 | air heater SCR | 4-20mA | NI9265_2 (2..2) | AO1-COM1 | N/A
O:1.6 | GLY_SCR | 156 | glycol heater SCR | 4-20mA | NI9265_2 (3..1) | AO2-COM2 | N/A
| OUTPUTS: | # | Module (used..left) |  | INPUTS: | # | Module (used..left)
| ON-OFF 24VDC | 4 | NI9403 (4..28) |  | IN 24VDC | 8 | NI9403 (8..15) |  | MOD5
| ON-OFF 120VAC | 5 | NI9403 (5..23) |  | IN 120VAC | 2 | NI9403 (2..13)
| ANALOG 4-20mA | 7 | NI9265 (7..8) |  | 4-20mA | 11 | NI9208 (11..5)
[5 empty] | T-type TCs | 2 | NI9211 (2..2)
[5 empty] | frequency | 1 | NI9403 (1..12)
INPUTS:
PLC |  | WIRE | DESC | SIGNAL | cDAQ-9174 (used..left) | Contact | RELAY | COUNTS (this kind.. Total) | COMMENTS
I:0/0 | GLY_TEMP_LSL | 10000 | GLY tank low level switch | 5VDC | NI9403 (8..24) | DIO.7 | N/A | N/A
I:0/1 | GLY_HEAT_OT | 10001 | GLY heater over temp | 5VDC | NI9403 (9..23) | DIO.8 | N/A | N/A | From High limit control
I:0/2 | AIR_HEAT_OT | 10002 | AIR heater over temp | 5VDC | NI9403 (10..22) | DIO.9 | N/A | N/A | From High limit control
I:0/3 | BLOWER_ON | 10003 | Blower on | 5VDC | NI9403 (11..21) | DIO.10 | N/A | N/A | From VSD
I:0/4 | PUMP_ON | 10004 | Pump on | 5VDC | NI9403 (12..20) | DIO.11 | N/A | N/A | From VSD
I:0/5 | COMP_OL | 10005 | Compressor over limit | 5VDC | NI9403 (13..19) | DIO.12 | N/A | N/A
I:0/6 | COMP_DIS_PSH | 10006 | Compressor Discharge pressure | 5VDC | NI9403 (14..18) | DIO.13 | N/A | N/A
I:0/9 | EMO_CLR | 10009 | Emergency OFF button clear | 120VAC | NI9403 (15..17) | DIO.14 | IAC5 (120VAC to 5VDC) | 1..8
I:0/10 | RESET_PB | 10010 | Base reset after EMO | 5VDC | NI9403 (16..16) | DIO.15 | N/A
I:0/11 | MASTER_SW | 10011 | Master power switch | 120VAC | NI9403 (17..15) | DIO.16 | IAC5 (120VAC to 5VDC) | 2..9
SENSORS:
PLC |  | WIRE | DESC | SIGNAL | cDAQ-9174 (used..left) | Contact | RELAY | COUNTS (this kind.. Total) | COMMENTS
I:2.0 | RH product temp | 10200 | Temperature -40 to 176 °F | 4 - 20mA | NI9208 (1..15) | AI0/COM | N/A | N/A | 251 is reference
I:2.1 | RH product | 10201 | HUM 0~100% | 4 - 20mA | NI9208 (2..14) | AI1/COM | N/A | N/A | 251 is reference
I:2.2 | RH nozzle temp | 10202 | Temperature -40 to 176 °F | 4 - 20mA | NI9208 (3..13) | AI2/COM | N/A | N/A | 251 is reference
I:2.3 | RH nozzle | 10203 | HUM 0~100% | 4 - 20mA | NI9208 (4..12) | AI3/COM | N/A | N/A | 251 is reference
I:2.4 | Ref suct pressure | 10204 | 0-200PSIG | 4 - 20mA | NI9208 (5..11) | AI4/COM | N/A | N/A | 251 is reference
I:2.5 | SPARE | 10205 | SPARE | 4 - 20mA | NI9208 (6..10) | AI5/COM | N/A | N/A | 251 is reference
I:2.6 | Comp suct temp | 301TE | Compressor suction temp | thermocouple | NI9211 (1..3) | TC0+/TC0- | N/A | N/A | T-Type thermocouple
I:2.7 | Comp disch temp | 302TE | Compressor discharge temp | thermocouple | NI9211 (2..2) | TC1+/TC1- | N/A | N/A | T-Type thermocouple
I:3.0 | Baro Pressure | 10300 | PT302 12.777 ~ 15.717 psia | 4 - 20mA | NI9208 (7..9) | AI6/COM | N/A | N/A | 251 is reference
I:3.1 | Nozzle dP | 10301 | Nozzle dP 0 to 5 inWC | 4 - 20mA | NI9208 (8..8) | AI7/COM | N/A | N/A | 251 is reference
I:3.2 | Air dP | 10302 | Air dP 0 to 0.1 inWC | 4 - 20mA | NI9208 (9..7) | AI8/COM | N/A | N/A | 251 is reference
I:3.3 | Glycol dP | 10303 | GLY dP 0 to 25psid | 4 - 20mA | NI9208 (10..6) | AI9/COM | N/A | N/A | 251 is reference
I:3.4 | GLY density | 10304 | GLY density | 4-20mA | NI9208 (11..5) | AI10/COM | N/A | N/A | 251 is reference
I:3.5 | GLY flow | 10305 | mdot 0 to 70lbm/min/ frequency out | frequency | NI9403 (18..14) | DIO.17 | N/A | N/A | 10355 is reference
I:3.6 | GLY inlet Pressure | 10306 | Product inlet Pressure 0-250psig | 4 - 20mA | NI9208 (12..4) | AI11/COM | N/A | N/A | Reflects new SETRA pressure transducers
I:3.7 | GLY outlet Pressure | 10307 | Product inlet Pressure 0-100psig | 4 - 20mA | NI9208 (13..3) | AI12/COM | N/A | N/A | Reflects new SETRA pressure transducers
|  |  | Compressor suction temp | thermocouple | NI9211 (1..3) | TC0+/TC0-
|  |  | Compressor discharge temp | thermocouple | NI9211 (2..2) | TC1+/TC1-
|  |  | Blower speed control | 4-20mA | NI9265_1 (1..3) | AO0-COM0
|  |  | Pump speed control | 4-20mA | NI9265_1 (2..2) | AO1-COM1
|  |  | Bypass damper | 4-20mA | NI9265_1 (3..1) | AO2-COM2
|  |  | Ref expansion valve AIR evap | 4-20mA | NI9265_1 (4..0) | AO3-COM3
|  |  | Ref expansion valve GLY evap | 4-20mA | NI9265_2 (1..3) | AO0-COM0
|  |  | air heater SCR | 4-20mA | NI9265_2 (2..2) | AO1-COM1
|  |  | glycol heater SCR | 4-20mA | NI9265_2 (3..1) | AO2-COM2
Note on E2: Carlos Herrera: Updated voltage once all is changed to REV C dwg status
Note on E27: Carlos Herrera: Updated voltage once all is changed to REV C dwg status
Note on E40: Carlos Herrera: Updated voltage once all is changed to REV C dwg status

## Visual_cDAQ-9174
[5 empty] | NI MAX |  |  | PHYSICAL LOCATION
| DESC | cDAQ-9174 (used..left) | WIRE | Contact | Task Name | Channel Name | Order in task | RELAY | I/O rack position |  | NI 9211 TC MOD 4 |  | NI 9211 TC MOD 5 |  | NI 9403 32ch Digital I/O MOD 6 |  |  | NI 9265 4ch ±20m OM MOD7 |  |  | NI 9265 4ch ±20m Omv MOD8
OUTPUTS | Air blower RUN | NI9403 (1..31) | 0 | DIO.0 | HXCal1_OUT_allDIGITAL | AirBlowerRUN | 0 | 70-ODC5 / 5VDC ctrl .. 3-60VDC load | 0/0 | 480 VAC PANEL
| Air blower STOP |  | 1
| Gly Pump RUN | NI9403 (2..30) | 2 | DIO.1 | HXCal1_OUT_allDIGITAL | GlyPumpRUN | 1 | 70-ODC5 / 5VDC ctrl .. 3-60VDC load | 0/1
| Gly Pump STOP |  | 3
| Air heater CON | NI9403 (3..29) | 4 | DIO.2 | HXCal1_OUT_allDIGITAL | AirHtrCON | 2 | OAC5 (5VDC to 120VAC) | 0/2
| Gly heater CON | NI9403 (4..28) | 5 | DIO.3 | HXCal1_OUT_allDIGITAL | GlyHtrCON | 3 | OAC5 (5VDC to 120VAC) | 0/3
| Compressor CON | NI9403 (5..27) | 8 | DIO.4 | HXCal1_OUT_allDIGITAL | CompCON | 4 | OAC5 (5VDC to 120VAC) | 1/0
| Hot gas solenoid | NI9403 (6..26) | 9 | DIO.5 | HXCal1_OUT_allDIGITAL | HotGasSoln | 5 | OAC5 (5VDC to 120VAC) | 1/1
| Suction cool solenoid | NI9403 (7..25) | 10 | DIO.6 | HXCal1_OUT_allDIGITAL | SuctCoolSoln | 6 | OAC5 (5VDC to 120VAC) | 1/2
| Blower speed control | NI9265_1 (1..3) | 150 | AO0-COM0 | HXCal1_OUT_allANALOG | BlowerVFD | 0 | N/A | N/A
| Pump speed control | NI9265_1 (2..2) | 151 | AO1-COM1 | HXCal1_OUT_allANALOG | PumpVFD | 1 | N/A | N/A
| Bypass damper | NI9265_1 (3..1) | 152 | AO2-COM2 | HXCal1_OUT_allANALOG | BypassDamper | 2 | N/A | N/A
| Ref expansion valve AIR evap | NI9265_1 (4..0) | 153 | AO3-COM3 | HXCal1_OUT_allANALOG | AirHTR SCR | 3 | N/A | N/A
| Ref expansion valve GLY evap | NI9265_2 (1..3) | 154 | AO0-COM0 | HXCal1_OUT_allANALOG | EXVairEVAP | 4 | N/A | N/A
| air heater SCR | NI9265_2 (2..2) | 155 | AO1-COM1 | HXCal1_OUT_allANALOG | GlyHTR SCR | 5 | N/A | N/A
| glycol heater SCR | NI9265_2 (3..1) | 156 | AO2-COM2 | HXCal1_OUT_allANALOG | EXVglyEVAP | 6 | N/A | N/A
INPUTS | GLY tank low level switch | NI9403 (8..24) | 10000 | DIO.7 | HXCal1_IN_allDIGITAL | GLY-tank-low | 0 | IAC5 (120VAC to 5VDC) | 2/0
| GLY heater over temp | NI9403 (9..23) | 10001 | DIO.8 | HXCal1_IN_allDIGITAL | GLY-htr-overtemp | 1 | N/A | N/A |  |  |  | 2 | used |  | 17 | used |  | 4 | used |  | 3 | used
| AIR heater over temp | NI9403 (10..22) | 10002 | DIO.9 | HXCal1_IN_allDIGITAL | AIR-htr-overtemp | 2 | N/A | N/A |  |  |  | 2 | available |  | 15 | available |  | 0 | available |  | 1 | available
| Blower on | NI9403 (11..21) | 10003 | DIO.10 | HXCal1_IN_allDIGITAL | BLOWER-on | 3 | N/A | N/A
| Pump on | NI9403 (12..20) | 10004 | DIO.11 | HXCal1_IN_allDIGITAL | PUMP-on | 4 | N/A | N/A
| Compressor over limit | NI9403 (13..19) | 10005 | DIO.12 | HXCal1_IN_allDIGITAL | COMP-overload | 5 | N/A | N/A | I/O RACKS
| Compressor Discharge pressure | NI9403 (14..18) | 10006 | DIO.13 | HXCal1_IN_allDIGITAL | COMP-disch-pressure | 6 | IAC5 (120VAC to 5VDC) | 2/1
| Emergency OFF button clear | NI9403 (15..17) | 10009 | DIO.14 | HXCal1_IN_allDIGITAL | EMO-CLR | 7 | IAC5 (120VAC to 5VDC) | 2/2
| Base reset after EMO | NI9403 (16..16) | 10010 | DIO.15 | HXCal1_IN_allDIGITAL | EMO-reset | 8 | N/A | N/A
| Master power switch | NI9403 (17..15) | 10011 | DIO.16 | HXCal1_IN_allDIGITAL | MASTER-power-switch | 9 | IAC5 (120VAC to 5VDC) | 2/3
SENSOR INPUTS | HX_GlyIn_1 | NI9217 |  | Mod1/ai0 | HXCal2_IN_ANALOG | HX_GlyIn_1 | 0
| HX_GlyIn_2 | NI9217 |  | Mod1/ai1 | HXCal2_IN_ANALOG | HX_GlyIn_2 | 1
| HX_GlyOut_1 | NI9217 |  | Mod1/ai2 | HXCal2_IN_ANALOG | HX_GlyOut_1 | 2
| HX_GlyOut_2 | NI9217 |  | Mod1/ai3 | HXCal2_IN_ANALOG | HX_GlyOut_2 | 3
| HX_AirIn1 | NI9217 |  | Mod2/ai0 | HXCal2_IN_ANALOG | HX_AirIn1 | 4
| HX_AirIn2 | NI9217 |  | Mod2/ai1 | HXCal2_IN_ANALOG | HX_AirIn2 | 5
| HX_AirIn3 | NI9217 |  | Mod2/ai2 | HXCal2_IN_ANALOG | HX_AirIn3 | 6
| HX_AirIn4 | NI9217 |  | Mod2/ai3 | HXCal2_IN_ANALOG | HX_AirIn4 | 7
| HX_AirOut1 | NI9217 |  | Mod3/ai0 | HXCal2_IN_ANALOG | HX_AirOut1 | 8
| HX_AirOut2 | NI9217 |  | Mod3/ai1 | HXCal2_IN_ANALOG | HX_AirOut2 | 9
| HX_AirOut3 | NI9217 |  | Mod3/ai2 | HXCal2_IN_ANALOG | HX_AirOut3 | 10
| HX_AirOut4 | NI9217 |  | Mod3/ai3 | HXCal2_IN_ANALOG | HX_AirOut4 | 11 | [5 empty] | NI 9217 RTD IM | [9 empty] | NI 9203 8ch ±20mA Analog IM
| HX_AirOut5 | NI9217 |  | Mod4/ai0 | HXCal2_IN_ANALOG | HX_AirOut5 | 12
| HX_AirOut6 | NI9217 |  | Mod4/ai1 | HXCal2_IN_ANALOG | HX_AirOut6 | 13 |  |  | EXTERNAL DAQ BOX
| HX_AirOut7 | NI9217 |  | Mod4/ai2 | HXCal2_IN_ANALOG | HX_AirOut7 | 14
| HX_AirOut8 | NI9217 |  | Mod4/ai3 | HXCal2_IN_ANALOG | HX_AirOut8 | 15
| Nozzle dP 0 to 5 inWC | NI9227_0 | 10301 |  | HXCal2_IN_ANALOG |  |  | N/A | N/A
| Nozzle dP 0 to 2.5 inWC | NI9227_1 | 10302 |  | HXCal2_IN_ANALOG |  |  | N/A | N/A
| UUT dP 0 to 0.1 inWC | NI9227_2 | 10306 |  | HXCal2_IN_ANALOG |  |  | N/A | N/A
| Liq Pin pump 0-200psig | NI9227_3 | 10205 |  | HXCal2_IN_ANALOG |  |  | N/A | N/A
| MassFlow Gly | NI9203_2 (1..7) | BR-WH | Mod8/ai0 | HXCal2_IN_ANALOG | GLY MassFlow | 24 | N/A | N/A
| RHin Temperature -40 to 176 °F | NI9203_0 (1..7) | 10200 | Mod8/ai0 | HXCal2_IN_ANALOG | RH in Temp | 16 | N/A | N/A
| RHin 0~100% | NI9203_1 (2..6) | 10201 | Mod8/ai1 | HXCal2_IN_ANALOG | RH in | 17 | N/A | N/A
| RHout Temperature -40 to 176 °F | NI9203_2 (3..5) | 10202 | Mod8/ai2 | HXCal2_IN_ANALOG | RH out Temp | 18 | N/A | N/A
| RHout 0~100% | NI9203_3 (4..4) | 10203 | Mod8/ai3 | HXCal2_IN_ANALOG | RH out | 19 | N/A | N/A
| Ref Suct Pressure | NI9203_4 (5..3) | 10204 | Mod8/ai4 | HXCal2_IN_ANALOG | Ref Suct Presure | 20 | N/A | N/A
| PT302 12.777 ~ 15.717 psia | NI9203_5 (6..2) | 10300 | Mod8/ai5 | HXCal2_IN_ANALOG | Barom Presure | 21 | N/A | N/A
| Setra Liq dP (0-25psid) | NI9203_6 (7..1) | 10303 | Mod8/ai8 | HXCal2_IN_ANALOG | Nozzle dP 0-5inWC | 22 | N/A | N/A
| Setra Liq In Pressure (0-100psig) | NI9203_7 (8..0) | 10304 | Mod8/ai7 | HXCal2_IN_ANALOG | UUT dP 0-pt1inWC | 23 | N/A | N/A
| Density Gly | NI9203_2 (2..6) | WH-BK | Mod8/ai1 | HXCal2_IN_ANALOG | GLY Density | 25 | N/A | N/A |  |  | used | 4 | 4 | 4 |  | 4 | 16 used | 4 | used | 1 | used | 8 | used
| Compressor suction temp | NI9211 (1..3) | 301TE | Mod1/ai0 | HXCal1_IN_TC_Internal | Comp-suct-temp | 0 | N/A | N/A |  |  | available | 0 | 0 | 0 |  | 0 | 0 available | 0 | available | NA | available | 0 | available
| Compressor discharge temp | NI9211 (2..2) | 302TE | Mod1/ai1 | HXCal1_IN_TC_Internal | Comp-disch-temp | 1 | N/A | N/A
| UUT dP 0 to 2.5 inWC | NI9203 | 10305 | Mod4/ai0 | HXCal2_IN_ANALOG
| Pump dP | NI9203 | 10307 | Mod4/ai1 | HXCal2_IN_ANALOG
0
1
2
#5
NI 9213 TC
#6
NI 9227 mA In
#6
NI 9411 D In

## Flow Data
PGLY 40%
Pump | Oberdorfer N991 |  |  |  |  | TEMP (°F) | Density (lbm/ft³) | µ (lbm/ft*hr) | Kinematic visco (cSt) | Cp (BTU/ lbm °F) | No correction factor needed because PGLY dynamic viscosity is <5 cSt
µ_Water @70°F |  | 1 | cSt |  |  | 50 | 65.89 | 17.465 | 6.8402 | 0.852
[6 empty] | 104 | 64.91 | 2.24 | 0.89055 | 0.878
[6 empty] | 100 | 64.9826 | 3.3678 | 1.3313 | 0.87607
Q(gpm) | 2 | 0.25 | 1.1 | 1.1
Q(ft³/hr) | 16.0428 | 2.0053 | 8.8235 | 8.8235
mdot (lbm/hr) | 1042.5015 | 130.3127 | 573.3758 | 573.3758
mdot (lbm/min) | 17.375
| 3/8 tube | 3/16 tube | 3/8 tube | 3/16 tube
Dh (ft) | 0.03125 | 0.015625 | 0.03125 | 0.015625
Area tube (ft²) | 0.00076699 | 0.00019175 | 0.00076699 | 0.00019175
V (ft/hr) | 20916.5341 | 10458.2671 | 11504.0938 | 46016.3751
Re | 12612.2727 | 3153.0682 | 6936.75 | 13873.5
Pressure drop in a pipe
Coil examined: | SF 12in BPO | 3W-?D-L
Temp examined (°F) | 100 | 100 |  |  | 0.0015
L (ft) | 26.6667 | 61.8 |  |  | 0.048
D (ft) | 0.03125 | 0.011
v (ft/s) | 348.6089 | 174.3045
g (ft/s²) | 32.174 | 32.174
rho (lb/ft³) | 64.9826 | 64.9826
f (Re > 4000)
f (Re <2000) | 0.035556 | 0.035556 |  | 0.032
dP Turbulent(psi)
dP Laminar(psi), 0.25gpm |  | 42561.7722
dP (per ft) Hazen-Williams | 0.27544 | 0.94983
dP for L Hazen-Williams | 7.3451 | 58.6997
mdot (lbm/hr) | 1042.5015 | 130.3127 | 573.3758 | 573.3758
Uncert (lbm/hr) | 0.001 | 0.001 | 0.001 | 0.001
Inlet Temp (°F) | 100
Uncert (°F) | 3.2
Outlet Temp (°F)
Uncert (°F)
CoilDesigner
Condenser | 3W_HD_WF_PGLY_asPlainFin.chx
Inlet Pressure given | 150 | psi
Inlet Temp given | 100 | °F
Q(gpm) | 0.385
Q(ft³/hr) | 3.0882
mdot (lbm/hr) | 200.6815
Mass flow given | 200
The pump is an Oberdorfer N991 we purchased from San Diego Pump; see attached spec sheet. It has a 1/2hp 1140rpm motor capable of 1.4 gpm at 0 psi, and 0.35 gpm at 150 psi. Switching to a 1750 rpm motor would give you about 1.2 gpm at 150 psi. Our system components are standard pipe fittings and refrigeration parts, so the system can handle the full 150 psi.
The viscosity of the fluid affects all the operational parameters of the pump: total head, flow and efficiency. The viscosity of water at 70 ºF is 1 cSt (centiStoke). If the viscosity of the fluid is 5 cSt or more a correction for viscosity has to be applied.
THE INFLUENCE OF SPECIFIC GRAVITY ON TOTAL HEAD (SpecificGravityOnTotalHead.pdf)

## Circulation Heater Calcs
2012-02-19
DATA FOR WATER FLOW
Copper tube
OD | 0.5 | in
wall thick | 0.032 | in
ID | 0.436 | in | 0.036333 | ft
mdot | 1000 | lbm/hr
Area | 0.0010368 | ft²
TEMP (°F) | Density (lbm/ft³) | µ (lbm/ft*s) | vel (ft/sec) | vel (ft/hr) | vol (ft³/hr) | gpm | Re
70 | 62.3013 | 0.00065512 | 4.3003 | 15481.1185 | 16.051 | 2.0012 | 14858.7933
150 | 61.1949 | 0.00028859 | 4.3781 | 15761.0223 | 16.3412 | 2.0373 | 33730.6053
200 | 60.1207 | 0.00020333 | 4.4563 | 16042.6215 | 16.6332 | 2.0737 | 47872.9907
For Circulation heaters from watlow (no issues)